# Medical Temperature Monitoring System

### ATmega328P / Arduino Uno — Embedded Systems Project

Platform: Arduino Uno / ATmega328P
Development & Simulation: Proteus VSM / VSM Studio
Sensor: LM35
Display: 16×2 LCD
Non-volatile Storage: Internal EEPROM
Alarm: LED + Buzzer
User Input: Three push buttons

## 1. Project Overview

The Medical Temperature Monitoring System is an embedded temperature-monitoring prototype designed around the ATmega328P microcontroller.

The project goes beyond basic temperature measurement. It focuses on the behavior of a complete embedded device, including system startup, self-testing, temperature classification, alarm management, measurement history, unit conversion, non-volatile data storage, watchdog protection, reset handling, standby behavior, and fault detection.

The system was developed and functionally simulated in Proteus with an emphasis on predictable state-based behavior and user-oriented operation.

The main design objective was to keep internal implementation details hidden from the end user and present only meaningful device information through the LCD and user interface.

# 2. System Architecture

The system is built around an ATmega328P / Arduino Uno.

### Main functional blocks

- LM35 temperature sensor

- 16×2 LCD user interface

- Measurement button

- Memory button

- Celsius/Fahrenheit button

- Fever LED indicator

- Audible buzzer

- EEPROM measurement storage

- Timer1 system timing

- Watchdog protection

- Startup / POST subsystem

- Reset handling

- Standby management

The firmware uses a state-based architecture to separate startup, self-test, ready, measurement, memory, unit-selection, reset, and failure states.

# 3. Hardware Interface

The current hardware mapping is:

| Function | ATmega328P Pin |
| --- | --- |
| LM35 output | A0 |
| LCD RS | D7 |
| LCD Enable | D6 |
| LCD D4 | D5 |
| LCD D5 | D4 |
| LCD D6 | D3 |
| LCD D7 | D2 |
| Fever LED | D8 |
| Buzzer | D9 |
| Measure Button | D10 |
| Memory Button | D11 |
| Unit Button | D12 |

The three buttons use the ATmega328P internal pull-up configuration and are activated by connecting the corresponding input to ground.

# 4. Temperature Measurement

Temperature is measured using the LM35 analog temperature sensor.

The sensor voltage is converted through the ATmega328P ADC and interpreted as temperature in Celsius.

The measurement chain is:

LM35 → ADC → Voltage Conversion → Temperature Calculation → Classification → User Display

The LM35 provides a linear temperature output, allowing the firmware to convert the ADC reading into degrees Celsius.

The current measurement logic is based on:

- ADC resolution: 10-bit

- Reference assumption: 5 V

- LM35 scale: 10 mV/°C

The resulting temperature is maintained internally in Celsius.

# 5. Temperature Classification

The system classifies measured temperature into three user-oriented conditions.

### NORMAL

Temperature < 37.5°C

### ELEVATED

37.5°C – 37.9°C

### FEVER

Temperature ≥ 38.0°C

Importantly, the classification logic remains based on Celsius thresholds even when the user selects Fahrenheit display.

This prevents unit conversion from affecting the actual medical classification thresholds.

# 6. Fever Alarm

When the measured temperature reaches the fever threshold, the system activates:

- Visual LED indication

- Audible buzzer

The alarm is associated with the current measurement event rather than continuously recalculating the alarm state in the background.

After the measurement interaction is completed and the measurement button is released, the alarm is explicitly turned off.

This provides a predictable user interaction rather than leaving the alarm continuously active.

# 7. Power-On Self-Test (POST)

A major feature of the project is the Power-On Self-Test.

At startup, the system does not immediately enter normal operation.

Instead, it performs a sequence of functional checks.

### POST sequence

- LCD interface check

- Temperature sensor plausibility check

- EEPROM / memory integrity check

- Alarm output test

- Button state check

- Timer functionality test

The user sees meaningful system information rather than internal diagnostic values.

Example:

SYSTEM CHECK

LCD          OK

followed by the other system checks.

When all checks pass:

SYSTEM READY

ALL SYSTEMS OK

If a critical POST condition fails:

SYSTEM FAILURE

SERVICE REQUIRED

# 8. Sensor Fault Detection

The temperature sensor is not blindly trusted.

During POST, the firmware performs a plausibility check on the sensor reading.

The expected operating range used by the current prototype is:

−10°C to +60°C

A value outside this range is treated as an abnormal sensor condition.

One of the functional tests performed during development was intentionally disconnecting the LM35 sensor.

The system detected the abnormal condition and entered the failure state instead of treating the invalid reading as a normal temperature measurement.

This test demonstrates that the system includes basic fault handling rather than simply displaying raw ADC data.

# 9. EEPROM Measurement Memory

The system provides non-volatile storage for five measurement records.

The memory is designed as a circular five-slot history.

The records are numbered:

1 → 2 → 3 → 4 → 5

After all five slots are occupied, the next measurement replaces record #1, then #2, and so on.

This prevents unlimited memory growth while maintaining a predictable fixed-size history.

Each record stores:

- Record number

- Temperature

- Unit used at measurement time

The stored temperature remains internally represented in Celsius, while the unit flag records whether the user was using Celsius or Fahrenheit.

# 10. Measurement History

When the Memory button is pressed, the stored records are presented to the user sequentially.

Example:

MEM #1

36.8°C

followed by:

MEM #2

98.2°F

The unit associated with each measurement is preserved.

This means an old measurement does not change its displayed unit merely because the user later changes the current system unit.

If there are no stored measurements, the system displays:

MEMORY EMPTY

# 11. Celsius / Fahrenheit Support

The Unit button allows the user to switch between Celsius and Fahrenheit.

The conversion is performed using the standard relationship:

°F = °C × 9/5 + 32

The system maintains Celsius as the internal reference for temperature classification.

This ensures that the fever thresholds remain consistent regardless of the selected display unit.

The selected unit is also stored in EEPROM so that the system can preserve the user's configuration.

# 12. Last Measurement Recovery

A design improvement was made to prevent the last measurement from being lost after a reset or restart.

Instead of relying only on a RAM variable, the firmware reconstructs the latest measurement from the EEPROM history.

This provides a single persistent source of truth for measurement history.

If records exist, the firmware determines the latest record according to the current memory index and record count.

This allows the last valid measurement to be restored after restarting the system.

# 13. Timer Architecture

Application timing is implemented using the ATmega328P Timer1 peripheral.

The firmware does not use:

- delay()

- millis()

Instead, Timer1 operates in CTC mode.

### Timer configuration

- CPU frequency: 16 MHz

- Timer1 prescaler: 64

- Compare value: 249

- Timer interrupt period: approximately 1 ms

The Timer1 interrupt maintains a system time counter.

Application functions use elapsed-time calculations based on this system timer.

This allows the firmware to perform timing without blocking the main application flow.

# 14. Non-Blocking Firmware Design

The project uses a state-machine-oriented architecture.

Major system states include:

- RESET

- STARTUP

- POST

- READY

- MEASUREMENT

- MEASUREMENT FINISH

- MEMORY

- UNIT

- POST FAILURE

This architecture allows the system to perform timed operations without blocking the main loop.

For example, POST steps advance according to Timer1 timing rather than using blocking delays.

This approach makes the firmware easier to extend and provides more predictable control over user interaction.

# 15. Button Handling

The three user buttons use software debouncing.

Each button has separate tracking for:

- Raw state

- Stable state

- State-change time

- Press event

A button press is converted into a one-shot event that can be consumed by the application.

This prevents mechanical switch bounce from being interpreted as multiple user commands.

The design also avoids exposing button-state details to the user interface.

# 16. Watchdog Protection

The ATmega328P watchdog timer is enabled as an additional system-protection mechanism.

The watchdog timeout is configured to approximately four seconds.

The firmware periodically resets the watchdog during normal operation.

If the application becomes stuck and stops servicing the watchdog, the microcontroller can automatically reset.

This provides an additional layer of resilience against unexpected firmware lockups.

# 17. Reset Handling

Hardware reset behavior was explicitly designed as part of the user experience.

A hardware reset is detected and presented to the user through the LCD:

DEVICE RESET

PLEASE WAIT

The reset notification is intentionally kept out of the normal Virtual Terminal output.

This follows the design principle that the user should see meaningful device information rather than internal diagnostic information.

After the reset screen, the system returns through startup and POST before becoming ready for operation.

# 18. Startup Experience

The startup sequence is intentionally separated from normal operation.

The user first sees:

TEMP MONITOR

STARTING...

The device then performs its POST sequence.

After successful initialization:

SYSTEM READY

ALL SYSTEMS OK

The system is then placed into its normal ready state.

The startup process therefore gives the user confidence that the device has completed its initialization and checks before operation.

# 19. Standby Mode

The system includes a 30-second inactivity timeout for the ready state.

After approximately 30 seconds without user activity, the device enters standby.

The user-facing standby screen is:

STANDBY

PRESS BUTTON

This provides a simple power/interaction management behavior without exposing internal timer information.

# 20. User Interface Philosophy

One of the key design decisions was to treat the LCD and Virtual Terminal as user interfaces, rather than development/debug consoles.

The user should not need to see:

- ADC values

- Sensor voltage

- Timer ticks

- EEPROM addresses

- Internal state variables

- Watchdog counters

- Debounce states

- Internal flags

- Proteus/VSM implementation details

Instead, the interface communicates meaningful information such as:

- Temperature

- Temperature classification

- System status

- Memory records

- Unit selection

- Alarm condition

- System failure

- Reset status

- Standby state

This separation between internal implementation and user-facing behavior is an important part of the project's design philosophy.

# 21. Virtual Terminal

The Virtual Terminal is used as an additional user/service interface.

It is not intended to expose raw firmware debugging information.

For example, the memory history can be presented as:

===== MEMORY =====

MEM #1 = 36.8 C

MEM #2 = 98.2 F

==================

System status can also be communicated through concise messages.

The LCD remains the primary user-facing interface.

# 22. EEPROM Data Structure

The EEPROM layout was deliberately organized rather than relying on arbitrary addresses.

The persistent data includes:

- EEPROM initialization marker

- Measurement count

- Five measurement records

- Next memory index

- Current unit configuration

A magic value is used to determine whether the EEPROM contains initialized application data.

On first use, the system initializes the required EEPROM structure.

Subsequent startups load and validate the stored configuration and measurement history.

# 23. EEPROM Validation

The POST memory test validates the stored structure before normal operation.

The firmware checks:

- EEPROM initialization marker

- Measurement count

- Next record index

- Record numbers

- Temperature plausibility

- Unit flags

This helps prevent corrupted EEPROM data from being silently treated as valid measurement history.

# 24. Alarm Self-Test

The POST sequence includes a functional alarm test.

The LED/buzzer alarm is activated for a controlled period and then turned off.

This confirms that the firmware can command the alarm output during startup.

The test is intentionally brief so that the alarm does not become disruptive during normal startup.

# 25. Timer Self-Test

Timer1 itself is included in POST.

The firmware starts a timer interval and verifies that the expected elapsed time is reached.

This allows the application timing mechanism to be checked before normal operation.

Because the entire application timing architecture depends on Timer1, verifying it during POST provides useful system-level confidence.

# 26. Button Startup Check

The POST sequence checks the initial button condition.

The current implementation verifies that the buttons are released when the system starts.

This prevents the system from entering normal operation while a user input is already active.

This is a functional startup-state check rather than a complete electrical diagnostic of the button wiring.

# 27. LCD Test Limitation

The LCD POST test is intentionally described as an interface check.

The current hardware uses the LCD primarily as a write interface. Therefore, the firmware cannot independently prove that every LCD pixel, segment, or physical connection is functioning correctly through software alone.

The current test confirms that the firmware executes the LCD initialization and output sequence.

A true hardware-level LCD diagnostic would require additional hardware/readback capability.

This limitation is explicitly recognized rather than claiming a capability the hardware does not provide.

# 28. Button Diagnostic Limitation

The same principle applies to the push buttons.

Because the buttons use internal pull-ups, a disconnected button input can appear electrically similar to a released button.

Therefore, the current POST verifies the logical startup state but does not provide complete open-circuit or wiring-fault detection.

A more advanced diagnostic design could use additional circuitry or different input arrangements to distinguish:

- Button released

- Button pressed

- Open circuit

- Wiring fault

The current prototype deliberately keeps the hardware simple.

# 29. Testing Performed

Functional testing has been performed in Proteus.

Tests include:

### Startup Test

Verified:

TEMP MONITOR

STARTING...

followed by POST.

### Sensor Test

Normal LM35 operation produces a valid temperature reading.

### Sensor Fault Test

The LM35 connection was intentionally disconnected.

The system detected the abnormal sensor condition and reported:

SYSTEM FAILURE

SERVICE REQUIRED

### Fever Test

The simulated temperature was increased above the fever threshold.

The system activated the LED and buzzer.

### Memory Test

Measurements were stored and displayed using the five available memory slots.

### Unit Test

The system was switched between Celsius and Fahrenheit.

### Reset Test

Hardware reset behavior was tested, including the user-facing reset screen and subsequent startup sequence.

### Watchdog Test

The watchdog mechanism was independently tested to verify that the microcontroller can recover from an application lock condition.

# 30. Design Decisions

Several design decisions were made specifically to make the prototype behave more like a complete embedded product.

### Decision 1 — Celsius as internal reference

Temperature classification is always performed in Celsius.

### Decision 2 — Fixed five-record memory

A small circular history provides predictable memory usage.

### Decision 3 — EEPROM persistence

Important user data and measurements survive reset.

### Decision 4 — Timer1 instead of blocking delays

The application remains responsive while performing timed operations.

### Decision 5 — State-based architecture

Different system functions are isolated into explicit states.

### Decision 6 — Watchdog protection

The system has an independent recovery mechanism for firmware lockups.

### Decision 7 — User-oriented diagnostics

Internal implementation details are hidden from the normal interface.

### Decision 8 — Explicit fault handling

Abnormal sensor conditions are handled as system faults rather than being silently displayed as valid measurements.

# 31. Engineering Focus

The project demonstrates practical experience with several embedded-system concepts:

- ATmega328P microcontroller programming

- Embedded C/C++

- Analog sensor interfacing

- ADC processing

- LCD interfacing

- Digital I/O

- Push-button debouncing

- Interrupts

- Timer1 configuration

- Non-blocking firmware architecture

- State machines

- EEPROM data persistence

- Data validation

- Watchdog timers

- Hardware reset handling

- Fault detection

- Alarm control

- Embedded user-interface design

- Hardware/software integration

- Proteus/VSM simulation

- Functional system testing

# 32. Current Prototype Scope

This project is an engineering prototype and simulation-based embedded system.

It should not be interpreted as a certified clinical medical device.

The project demonstrates embedded-system architecture, firmware design, interface behavior, fault handling, and functional simulation.

A production medical device would require substantially more work, including appropriate sensor accuracy validation, electrical safety, EMC testing, environmental testing, risk management, verification and validation, regulatory compliance, and applicable medical-device standards.

# 33. Future Development

Possible future improvements include:

- More advanced sensor diagnostics

- Improved electrical fault detection for buttons

- Hardware-level LCD diagnostics

- Sensor calibration

- Measurement filtering

- Additional measurement history

- Battery monitoring

- Low-power operation

- More sophisticated alarm patterns

- Data export

- Real-time clock integration

- Event logging

- Production PCB design

- Hardware validation outside simulation

- Formal verification and validation

- Medical-device compliance analysis

# 34. Conclusion

The Medical Temperature Monitoring System was developed as a practical embedded-system project focused on reliability, structured firmware design, user interaction, and fault handling.

The project evolved from a basic temperature measurement concept into a more complete device architecture incorporating:

Measurement → Classification → Alarm → Memory → Self-Test → Fault Detection → Recovery

The most important engineering objective was to design the behavior of the complete system rather than simply demonstrate that an LM35 can be read by an Arduino.

The resulting prototype demonstrates how hardware interfaces, firmware architecture, persistent storage, timing, diagnostics, and user experience can be integrated into a single embedded system.

## Project Status

Prototype / Functional Simulation

Platform: ATmega328P / Arduino Uno
Simulation: Proteus VSM
Primary Sensor: LM35
User Interface: 16×2 LCD + Push Buttons
Storage: EEPROM
Timing: Timer1 Interrupt Architecture
Protection: Watchdog Timer
Testing: Functional Proteus Testing

### Key Message

This project is not just a temperature sensor demonstration. It is an embedded-system prototype designed around system behavior, reliability, user interaction, persistent data, and fault handling.